Laser Safety

Policy and Guidance

Based on AURPO Guidance Note No. 7 (2012)

‘Guidance on the Safe Use of Lasers in Education and Research’

Contents

Laser Safety Policy

- Policy Statement	…………………………………………………………………………..	4

- Scope	……………………………………………………………………………………	4

- Legislation		…………………………………………………………………………..	4

- Responsibilities	…………………………………………………………………………..	4

4.1.	Heads of Schools / Centre Directors	………………………………………..	4

4.2.	University Laser Safety Officer	………………………………………..	4

4.3.	Departmental Laser Safety Supervisor	……………………………….	5

4.4.	Researcher / Supervisor / Principal Investigator	………………………	5

4.5.	Laser Users	………………………………………………………………….	5

5.   Laser Equipment	…………………………………………………………………………..	5

6.   Laser Laboratories and Laser Controlled Areas	………………………………………..	6

7.   Risk Assessment	…………………………………………………………………………..	6

8.   Control Measures	…………………………………………………………………………..	6

8.1.	Use of Control Measures	…………………………………………………	6

8.2.	Maintenance, Examination and Testing of Control Measures	……..	7

9.   Monitoring Exposure	…………………………………………………………………..	7

10. Health Surveillance	…………………………………………………………………..	7

11. Information, Instruction and Training	…………………………………………………	7

12. Emergency Procedures	…………………………………………………………………..	8

13. Record Keeping	……………………………………………………………………………	8

14. Monitoring and Audit	………………………………………………………………….	8

Laser Safety Guidance

- Introduction	……………………………………………………………………………	9

- Hazard Classification for Lasers	………………………………………………………….	9

2.1.	Class 1	…………………………………………………………………..	9

2.2.	Class 1M	…………………………………………………………………..	10

2.3.	Class 2	…………………………………………………………………..	10

2.4.	Class 2M	…………………………………………………………………..	10

2.5.	Class 3R	…………………………………………………………………..	10

2.6.	Class 3B	…………………………………………………………………..	10

2.7. 	Class 4	…………………………………………………………………..	10

2.8.	Example AELs	………………………………………………………….	11

2.9.	Maximum Permissible Exposures	…………………………………………	11

- General Safety Precautions	…………………………………………………………..	11

3.1.	Laser Inventory	…………………………………………………………..	11

3.2.	Registration of Users	…………………………………………………………..	11

3.3.	Training	……………………………………………………………………	12

3.4.	Undergraduate Work	…………………………………………………………..	12

3.5.	Labelling of Lasers	…………………………………………………………...	13

3.6.	Designation of Laser Area	…………………………………………………..	13

3.7.	Laboratory Design	…………………………………………………………...	13

3.8.	Experimental Set-Up	…………………………………………………………..	14

- Precautions for specific classes of laser	…………………………………………………..	14

4.1.	Class 1	……………………………………………………………………	14

4.2. 	Class 1M and 2M	…………………………………………………………..	15

4.3.	Class 2	…………………………………………………………………...	15

4.4.	Class 3R	……………………………………………………………………	15

4.5.	Class 3B or 4	……………………………………………………………………	15

- Hazard / Risk Assessment	…………………………………………………………..	15

5.1.	Stages in a Risk Assessment	…………………………………………………..	16

5.2.	Assessing risk	…………………………………………………………………....	16

5.3.	Control Measures	……………………………………………………………16

5.4.	Laser Controls – Optical Hazards	………………………………………….	17

5.4.1	Laser Alignment	…………………………………………………………..	17

5.5.	Control of Non-Optical Hazards	…………………………………………	18

5.6.	Management Controls	………………………………………………….	18

5.7.	Laser Accident	………………………………………………………......	19

5.8.	Laser Incident Grab Sheet	…………………………………………………..	19

5.9.	Assessing residual risk and recording the results	………………………..	19

- Personal Protective Equipment	…………………………………………………………...	19

- Medical Supervision, Emergency Eye Examinations and Accidental Exposures ……..	20

- Laser Pointers	…………………………………………………………………………….	21

Appendices:

- Laser Inventory	…………………………………………………………………………….	22

- Registration form for Laser Users		…………………………………………………..	23

- Use of Class 1M, 2M and 3R Lasers Risk Assessment	…………………………………	24

- User of Class 3B and 4 Lasers Risk Assessment	…………………………………	27

- Example Scheme of Work with Procedure	………………………………………….	34

- Summary of Warning and Protective Control Measures	………………………………...	41

- Summary of BS EN 60825-1 Manufacturers and User Requirements	……………….	42

- The Safe Use of Laser Pointers	…………………………………………………………..	43

- Laser Incident Grab Sheet	…………………………………………………………..	44

- Laser Signs and Labels	……………………………………………………………………	45

- Optical Hazards and Biological Effects of Laser Radiation	………………………..	48

- Training Record Form	……………………………………………………………………	49

Laser Safety Policy

- Policy Statement

The University is committed to ensuring that exposure of its employees, students and visitors to Non-Ionising radiation from lasers is either prevented or, where that is not reasonably practicable, is reduced to an acceptably low level (below the Maximum Permissible Exposure Standard).

- Scope

This guidance applies to all staff, students and visitors working with lasers on Queens University premises and in particular to those staff members who have managerial responsibilities placed upon them to actively monitor the implementation of this guidance.

- Legislation

The Control of Artificial Optical Radiation at Work (Northern Ireland) Regulations came into force in 2010 with the aim to protect workers from the risks to health from hazardous sources of artificial optical radiation (AOR) including lasers. The Regulations include the requirement to assess and control the hazards from light emitted from all artificial sources in all its forms such as ultraviolet, infrared and laser beams.

- Responsibilities

4.1	Heads of Schools / Centre Directors are responsible for:

- Ensuring that all work with lasers in their area of responsibility is carried out in accordance with the University Laser Safety Policy and with this Code of Practice.

- Appointing a Departmental Laser Safety Supervisor where Class 3B and 4 lasers are used.

- Adopting local rules for lasers in their area of responsibility, where Class 3B and 4 lasers are used.

4.2	University Laser Safety Officer (LSO) (University Safety Service) is responsible for ensuring that arrangements are in place for:

- The training of new staff / students;

- Identification of lasers and users of equipment;

- Provision of a measuring service (where appropriate);

- Inspection of all new laser facilities; and

- The routine auditing of laser facilities.

4.3	Laser Safety Supervisor (LSS) should ensure that:

- All lasers except for low power Class 1 devices (and excluding laser printers, DVDs, Class 2 laser pointers etc.) are identified;

- All lasers are labelled in accordance with appendix 10 and laser designated areas clearly identified;

- Schemes of work are drawn up, where necessary, for the safe operation of lasers (see example in Appendix 5). These will normally be required for all Class 3B and Class 4 lasers when not totally enclosed;

- Personnel intending to work with Class 3R, Class 3B and Class 4 lasers, and others who may be working with modified Class 1M or Class 2M devices, will need to be identified and receive training in the safe use of lasers;

- Laser safety eyewear are provided and worn (where appropriate) by all people working with Class 3B and Class 4 lasers when the beam is not totally enclosed and that training is given in the use and maintenance of this eyewear;

- Undergraduates working with lasers should use the minimum power laser practicable and follow a written scheme of work;

- All lasers in the department are used in accordance with this guidance; and

- Routine surveys are undertaken to ensure compliance with this guidance.

4.4	Researcher / Supervisor / Principal Investigator is responsible for ensuring that:

- All work involving hazardous lasers are covered by risk assessments and where appropriate, written schemes of work and protocols.

- Procedures are in place to guarantee that lasers are made safe prior to disposal and dealt with appropriately if they contain hazardous materials;

- Laser workers are effectively trained in the operating techniques required and that inexperienced staff are adequately supervised.

4.5	Laser Users are responsible for:

- Observing the Policy / Guidance and Schemes of Work applicable to the lasers that will be used an to follow the guidance of supervisors and the Laser Safety Supervisor;

- Not leaving a laser experiment running unattended unless a risk assessment has established that it is safe to do so;

- Their own safety and that of others who may be affected by their acts or omissions; and

- Wearing the appropriate laser eyewear where required when working with Class 3B or Class 4 lasers where there is the possibility of stray laser beams that could damage the eyesight.

- Laser Equipment

Only laser equipment which has been manufactured, classified, labelled and in conformance with the standard BS EN 60825-1:2007 Safety of Laser Products – Part 1: Equipment classification and requirements ‘Safety of Laser Products’ may be brought onto and used in University premises.

The Laser Safety Supervisor shall maintain an inventory of all laser products greater than Class 2 held within the areas under their direct control. This inventory shall be updated as necessary and forwarded to the University Laser Safety Officer.

Only Class 1, 2 or visible beam Class 3R laser products shall be used for demonstration, display or entertainment purposes in unsupervised areas.

- Laser Laboratories and Laser Controlled Areas

No work with laser products of higher Class than 2 shall take place on any University premises or off-site, until the areas have been inspected and approved by the Laser Safety Officer.

Such work areas shall be physically demarcated or, where this is not reasonably practicable, delineated from other working areas. Entrances to these areas must be clearly labelled with ‘Laser Safety’ signs and doors (where applicable) safety interlocked. Access shall be restricted to authorised personnel only.

- Risk Assessment

No work, which is liable to expose and employee, student and/or visitors to laser radiation in excess of the applicable maximum permissible exposure (MPE) shall be carried out until a suitable and sufficient risk assessment has been made.

The person(s) directly involved in the work, the Laser Safety Supervisor and, so far as is reasonably practicable, employee safety representatives shall participate in the risk assessment process. Advice can be sought from the University Laser Safety Officer if required.

The Laser Safety Supervisor shall maintain a register of risk assessments and review them periodically to ensure that they remain suitable and sufficient.

- Control Measures

Wherever reasonably practicable, laser hazards should be eliminated completely at source by the use of engineering controls (e.g. by total enclosure of the beam). The feasibility of using a laser of a lower class must also be considered in controlling hazards. Where neither of these options is reasonably practicable, control measures must be considered and applied in the following order of priority.

- Engineering controls, such as barriers and guards to prevent access to laser radiation.

- Administrative controls such as schemes of work, warning signs and training.

- The use of Personal Protective Equipment (PPE) e.g. laser safety goggles.

8.1	Use of Control Measures

The workplace shall be inspected on a regular basis to ensure that the control measures recommended in the risk assessment are in place and are being properly used and / or applied.

8.2	Maintenance, Examination and Testing of Control Measures

All active engineering controls and design features (e.g. interlocks, shutters), and personal protective equipment shall be maintained in an efficient state, in efficient working order, in good repair and in a clean condition.

Protective eyewear should be checked periodically for signs of wear or damage. The eyewear must be replaced when necessary. For example, protective eyewear which has been subjected to a single incident of accidental exposure at a high irradiance, should be replaced.

Protective eyewear should also be examined for suitability on each occasion prior to use.

- Monitoring Exposure

The Laser Safety Supervisor in consultation with the University Laser Safety Officer shall estimate the exposure to laser radiation in the workplace and at those locations at which it is reasonably foreseeable that a person might be present and where the highest levels of exposure could occur. This evaluation shall take into account intrabeam viewing and viewing of specular and diffuse reflections.

- Health Surveillance

Routine ophthalmic examinations of staff and students working with laser equipment have no value as part of a health surveillance programme. Ophthalmic examinations are sometimes carried out for other e.g. medico-legal reasons. Some of the investigative procedures used are themselves hazardous and these should therefore only be carried out when medically advisable and not used for routine screening.

- Information, Instruction and Training

The University will provide all employees and students who are engaged in work with lasers with such information, instruction and training as is necessary to enable them to conduct their work in accordance with the Regulations.

This information, instruction and training should include:

- The University’s policy for safe laser use;

- The risks of harm arising from the use and, reasonably foreseeable misuse of the laser equipment;

- The correct use and operation of the laser equipment, associated equipment and PPE;

- The meaning of displayed warning signs;

- The content of Schemes of Work and other working procedures;

- The procedures to be followed in the event of a suspected exposure to laser radiation.

All laser users will be informed by the Laser Safety Supervisor of the findings of risk assessments and local rules / schemes of work for their working area. Users must sign a copy of the local rules / Scheme of work where these apply and the signed copy shall be retained by the Laser Safety Supervisor.

- Emergency Procedures

Arrangements shall be in place to deal with accidents, incidents and emergencies involving laser systems.

The arrangements shall include:

- the procedure to be followed in the event of an ocular injury resulting from exposure to laser radiation;

- the provision of a ‘laser incident grab sheet’ for each laser system in the School / Centre;

- the provision of adequate first-aid facilities;

- the provision of fire-fighting equipment;

- Record Keeping

Records shall be kept by the Laser Safety Supervisor of relevant risk assessments, the laser equipment inventories, authorised laser personnel, of any incidents or accidents and of instruction and training given in relation to work with laser systems.

- Monitoring and Audit

Compliance with the provisions of this policy shall be monitored by the School / Directorate annually or more frequently if circumstances dictate. All of the appropriate records shall be made available on request for inspection by the audit team.

Laser Safety Guidance

- Introduction

The word laser is an acronym for Light Amplification by the Stimulated Emission of Radiation. The ‘light’ produced by a laser, a form of non-ionising optical radiation, has a unique combination of spatial coherence; monochromaticity and usually high collimation. This combination of characteristic distinguishes laser radiation from all other light sources.

Queens University uses a variety of lasers within its teaching, research, medicine and dentistry activities. Lasers emit radiation as narrow concentrated beams of light, not necessarily visible to the human eye. The optical and skin hazards presented by lasers vary markedly according to the wavelength and power of the output. The hazards of lasers are often associated with the ability of the laser to damage eyesight or burn skin, but quite often the radiation or optical hazards are not the ones that present the greatest risk, with associated risks from electrical supplies, cryogenic liquids or chemical dyes being more hazardous.

This guidance details the management arrangements and responsibilities in place for the safe use of lasers within the University.

All persons working with lasers (Class dependent) are subject to potential risks of exposure which may, at certain levels, be extremely harmful. The University is committed to protecting the health and safety of all staff, students, visitors and others who may be affected by its work.

- Hazard Classification for Lasers

Lasers are grouped into classes to give an indication of their potential to cause harm.  The higher the number in the classification series, the greater is the potential for the laser to cause harm.  The classification of a laser is determined by its Accessible Emission Limit (AEL).  This parameter is the maximum level of laser radiation which the laser can emit over its full range of capability during operation at any time after its manufacture.  The AEL is further defined by the wavelength, energy content and pulse characteristics (continuous wave or pulsed) of the laser beam.

There are seven classes of lasers: 1, 1M, 2, 2M, 3R, 3B and 4.

2.1	Class 1: Lasers that are safe under reasonably foreseeable conditions of operation, either because of the inherently low emission of the laser itself, or because of its engineering design such that it is totally enclosed and human access to higher levels is not possible under normal operation. NB if access panels of a totally enclosed system are removed for servicing etc. then the laser product is no longer Class 1 and the precautions applicable to the embedded laser must be applied until the panels are replaced.

2.2	Class 1M: Laser products emitting in the wavelength range 302.5nm to 4000nm, whose total output is in excess of that normally permitted for Class 1 laser products but because of their diverging beams or very low power density do not pose a hazard in normal use and comply with the measurement conditions for a Class 1M product. However, they may be hazardous to the eyes under certain conditions if gathering optics are used with them i.e.

- With a diverging beam if optics are placed within 100mm of the source to concentrate / collimate the beam.

- With a large diameter collimated beam viewed with binoculars or a telescope.

2.3	Class 2: Lasers that only emit visible radiation in the wavelength range from 400nm to 700nm, whose output is less than the appropriate AEL. They are safe for accidental viewing as eye protection is afforded by aversion responses, including the blink reflex. This reaction may be expected to provide adequate protection under reasonably foreseeable conditions of operation including the use of optical instruments for intrabeam viewing.

2.4	Class 2M: Laser products that only emit visible radiation in the wavelength range 400nm – 700nm, whose total output is in excess of that normally permitted for Class 2 laser products but because of their diverging beams or very low power density are safe for accidental viewing during normal use and comply with the measurement conditions for a Class 2M product. However they may be hazardous to the eyes under certain conditions if gathering optics are used with the, i.e.

- With a diverging beam if optics are placed within 100mm of the source to concentrate / collimate the beam.

- With a large diameter collimated beam viewed with binoculars or a telescope.

2.5	Class 3R: Lasers that emit in the wavelength range from 180nm – 1mm where direct intrabeam viewing is potentially hazardous but the risk is lower than for Class 3B lasers, and fewer manufacturing requirements and control measures for the user apply. The AEL is restricted to no more than five times the AEL of Class 2 for visible wavelengths and no more than five times the AEL of Class 1 for other wavelengths.

2.6	Class 3B: Lasers that are normally hazardous when direct intrabeam exposure occurs (i.e. within the Nominal Ocular Hazard Distance (NOHD), which is the distance within which the beam irradiance or radiant exposure will exceed the appropriate MPE). Viewing diffuse reflections is normally safe. Output levels must be less than the appropriate AELs for Class 3B devices.

2.7	Class 4: High power lasers that exceed the AELs for Class 3B products that are also capable of producing hazardous diffuse reflections. They may cause skin injuries, could also constitute a fire hazard and could cause hazardous fumes to be produced as well as being a hazard to the eyes. Their use requires extreme caution.

2.8	Example AELs

The AELs for He-Ne lasers emitting a narrow beam in CW mode at 633nm are as follows:-

- Class 1 and 1M		0.39mW

- Class 2 and 2M		1mW

- Class 3R			5mW

- Class 3B			500mW

These limits will also apply to other narrow beam CW lasers operating in the wavelength range 400 – 700nm except for Class 1 and 1M devices where there are further restrictions for wavelengths < 500nm.

2.9	Maximum Permissible Exposures

The MPE is that level of laser radiation to which, under normal circumstances, persons may be exposed without suffering adverse effects. The MPE levels represent the maximum level to which eye or skin can be exposed without consequential injury immediately or after a long time and are related to the wavelength of the radiation, the pulse duration or exposure time, tissue at risk and, for visible and near infra-red radiation in the range of 400nm – 1400nm, the size of the retinal image. MPE levels are specified in Safety of Laser Products – Part 14: A user’s guide, PD IEC/TR 60825-14.

The main criterion therefore for assessing the optical safety of a given laser in a given set of operating circumstances it the MPE.

Planned ocular exposure to laser radiation should not exceed the MPE for the intended exposure duration. However, for visible laser beams, sub-damage threshold effects, such as distraction, dazzle, glare and afterimages also should be considered. Software packages such as Lasermet and LaserBee are available from the University Safety Service that can be used to assist in the calculation of MPEs.

- General Safety Precautions

3.1	Laser Inventory

As a general rule, all Class 3R, 3B and Class 4 devices should be identified and a laser inventory for each School / Centre maintained (Appendix 1). The Laser Safety Supervisor will maintain these lists and notify the Laser Safety Officer of any changes. Before the operation of any new laser or significant modifications the Laser Safety Supervisor / Laser Safety Officer should be consulted.

3.2	Registration of Users

In order to make people aware of the hazards of lasers and to ensure that safe systems of work are being practiced, management arrangements should be in place to identify users of lasers. Risk assessments should also identify users of lasers and a list / record must be kept up to date in each School / Centre.

All persons intending to work with Class 3R, 3B and Class 4 lasers are required to register as a user by completing the registration form in Appendix 2 and forwarding to the Laser Safety Officer.

  - Training

Initial training will be a basic instruction in laser hazards, risks and their control. Class 3R, 3B and Class 4 laser workers must attend training before commencing any laser work and should also be familiar with the schemes of work / protocols required.

Schools / Centres are responsible for ensuring that their staff and students are adequately trained to carry out their work safely. Training in the use of individual lasers is the responsibility of the Research Supervisor / Principal Investigator and a record of this training must be made. The degree of training required by an individual laser user will be dependent on the amount of relevant experience and the nature of the work to be undertaken. As a minimum, the laser user will be required to work through the “Limits” interactive training system, view the Laser Safety Training modules available online at the University Safety Services website and then complete a laser safety training questionnaire.  The pass mark of 80% must be attained before the user is allowed to commence laser work.  In addition the School will provide specific training on the laser equipment, local rules and “in-house” emergency procedures.

Laser workers should be sufficiently competent in the operation and use of the equipment and should understand:

- The general nature of laser radiation;

- The health hazards, the tissues of the body at risk and the severity of harm which can result.

- The different laser classes and the meaning of the warning labels appropriate to the classes;

- The proper use of hazard control procedures and where appropriate the need for personal protective equipment;

- The need for any necessary additional precautions when undertaking non-routine activities; and

- Be familiar with the procedures and policy governing laser use, including emergency action and accident reporting procedures.

3.4	Undergraduate Work

If reasonably practicable, undergraduate work should be restricted to Class 1/1M, 2/2M or visible 3R lasers, especially for class experiments.

Clear written instructions should be provided for each student experiment and local rules should be drawn up and posted in the laboratory to introduce the students to good safety practice.

Students involved in project work and working with Class 3B or Class 4 lasers should be treated as laser workers and be subject to the normal registration and training process. The risk assessment should take into account the inexperience of the users and additional close supervision is likely to be necessary.

3.5	Labelling of Lasers

Lasers that are safe within reasonable foreseeable conditions of operation in Class 1 do not need warning labels. Supplementary information describing the laser product as a ‘Totally Enclosed System’ with details of the embedded laser clearly displayed is required in situations where access to the embedded product is routinely required. All other laser products should carry the appropriate warning labels in accordance with BS EN 60825-1 (Appendix 10).

3.6	Designation of Laser Area

A ‘laser controlled area’ is an area where the occupancy and activity of those within is subject to control and supervision for the purpose of protection from the laser radiation hazards. Access to a laser controlled area should be restricted to all except authorised persons when the laser is turned on and there is the possibility of exposure. Class 3B and 4 lasers should be operated in a laser controlled area if engineering controls have not entirely eliminated the risks; in this case an interlock system should be fitted (or other means of protecting unauthorised people, e.g. by controlling access by swipe card or key access).

The points of access to areas in which Class 3B and Class 4 lasers are used must be marked with appropriate warning signs (Appendix 10). There may also be experiments where open beam work with modified Class 1M/Class 2M or Class 3R lasers are used that will also warrant the display of appropriate warning signs.

3.7	Laboratory Design

The following considerations relate mainly to the use of Class 4 lasers but some may be appropriate for Class 3B devices as well, or as general specifications for a laser laboratory.

If practicable the laser laboratory should have a high level of illumination that will minimise pupil size and reduce the risk of stray laser light reaching the retina. Windows should be kept to a minimum and may need to be fireproof where higher-powered lasers are used.

Walls, ceilings and fittings should be painted with a light coloured matt paint to enhance illumination and minimise specular reflections. Reflecting surfaces such as the use of glass-fronted cupboards should be avoided.

Ventilation is important, especially with higher-powered lasers if cryogens are used, or if toxic fumes are produced that need to be extracted and in this case it is important that the extraction is very close to the source. Facilities may also be needed for the handling of toxic chemicals that are associated with some dye lasers.

The laboratory should be equipped with appropriate fire-fighting equipment. Electrical supplies and switch and control gear should be sited in order to:

- Enable the laser to be shut down by a person standing next to the laser;

- Enable the laser to be made safe in an emergency from outside the laser area if reasonably practicable;

- Prevent accidental firing of a laser;

- Provide an indication of the state of readiness of the laser;

- Enable personnel to stand in a safe place;

- Provide sufficient and adequate power supplies for all ancillary equipment and apparatus so that the use of trailing leads is minimised.

3.8	Experimental Set-up

Before starting to use your laser there are a number of basic risk reduction measures that should be considered:

- Can a lower powered laser be used?

- Can the output power of the laser be restricted if full power is not needed?

- Can intra-beam viewing be prevented by engineering design?

- Can the laser be used in a screened off area – limiting potential for others to be affected?

- Can work be carried out in a total enclosure?

- Beam paths should be as short as possible, optical reflections should be minimised and the beam terminated with an energy absorbing non-reflective beam stop.

- The laser and optical components should be securely fixed to the bench to avoid displacement and unintended beam paths.

- If practicable, align powerful lasers with low-power devices that are safe for accidental viewing, or reduce the power of the laser by turning down or introducing neutral density filters. The aim should be to get the output power <1mW, NB some kW lasers will only be able to be turned down to a few watts. Alternatively remote viewing techniques can be used.

- Eliminate chance of stray reflections – use coated optical components or shroud them so that only the intended beam can be refracted or reflected. Keep optical bench free from clutter and remove jewellery, wristwatches etc.

- And don’t forget to have the laser pointing away from the laboratory entrance!

- Precautions for specific classes of laser

4.1	Class 1 Laser products / systems do not require any special precautions or formal control measures. The exception is where such a system consists of a totally enclosed Class 3B/4 laser and access is required for the purpose of servicing or alignment. In this case, it should be included on the laser inventory and there should be a Scheme of Work for that activity.

4.2	Class 1M and 2M products can be hazardous if the output is viewed with optical instruments. Modification of these products needs to be carefully assessed, reclassification may be necessary and the appropriate control measures detailed.

4.3	For Class 2 laser products, protection is based on exposure being limited by the natural aversion response (0.25sec). Simple measures such as information to users not to deliberately aim the beam at people, stare into the beam and terminate the beam at the end of its useful path will be sufficient.

4.4	For Class 3R products, the control measures will include:

- Terminating the beam at the end of its useful path;

- Avoiding beam paths at eye level and where practicable enclose the beam;

- Instruction and training to an appropriate level;

- Taking care to prevent unintentional specular reflections;

- Where non-visible wavelengths are used an emission indicator device is to be used to indicate the laser is energised.

4.5	For Class 3B or 4 lasers, each laser laboratory or experiment, as appropriate, should have its own Laser Scheme of Work based upon the conclusions of a risk assessment.

The name of the Laser Safety Officer, Laser Safety Supervisor and the permitted authorised users, the extent of any laser designated area and reference to specific protocols that are to be used should be included in the scheme of work. (Example schemes of work can be found in Appendix 5). The Scheme of Work should be displayed in a prominent position or readily available within a laboratory folder.

All the above control measures indicated for Class 3R lasers should be used as well as remote interlocks, safety interlocks, key control, beam stop/shutter and the full list of user precautions detailed in Appendix 7. Areas need to be defined and warning signs used.

A summary of warnings and protective control measures from BS EN 60825 for all classes of laser is given in Appendix 7.

- Hazard / Risk Assessment

Excessive exposure to laser radiation will result in biological damage with the main areas of risk being the eyes and skin. Visible and near infra-red lasers are a special hazard to the eye because the very properties necessary for the eye to be an effective transducer of light result in high radiant exposure being presented to highly pigmented tissues. The biological effect of irradiation of skin by lasers operating in the visible and infra-red spectral regions may vary from a mild erythema to severe blisters.

When carrying out a risk assessment for the optical hazard, it is important to know of the effects of laser radiation on biological tissue.

Before the appropriate controls can be selected and implemented, laser hazards must be identified and evaluated together with non-beam hazards that may be present. The laser’s capability of injuring personnel and the environment and the way in which the laser or lasers are to be used needs consideration. A risk assessment must be carried out to establish the significant risks and whether suitable and effective controls exist.

5.1	Stages in a Risk Assessment

There are 5 stages to a risk assessment:

- Identify the hazards, i.e. the potential for harm;

- Assess risk from these hazards and who is at risk;

- Determine and implement the necessary control measures;

- Assess residual risk – repeating stage 3 if necessary; and

- Record the findings.

To assist in identifying hazards and risk control measures, a proforma has been developed for Class 1M, 2/2M and 3R lasers (Appendix 3) and Class 3B & 4 laser systems (Appendix 4).

Class 3B and Class 4 lasers are capable of causing significant eye injury to anyone who looks directly into the beam or its specular reflections. Diffuse reflections of a high-power laser beam can also cause permanent eye damage. High-power laser beams can burn exposed skin, ignite flammable materials and heat materials releasing hazardous fumes, gases or debris. Equipment and optical apparatus required to produce and control laser energy may also introduce additional hazards associated with high voltage, high pressure, cryogenics, noise, other forms of radiation, flammable materials and toxic fluid. Each proposed experiment or operation involving a laser must be evaluated to determine the hazards involved and the appropriate safety measures and controls required.

5.2	Assessing Risk

The people who may be at risk need to be identified. These may include cleaners, service personnel and other contractors, visitors, members of the public as well as trained operatives.

In laser safety it is usually more important to eliminate the risk of injury by adopting appropriate control measures in all situations where there is the possibility of MPEs being exceeded.

5.3	Control Measures

A risk control hierarchy must be applied to remove, or where his is not practicable, to reduce the level of risk. The hierarchy of control is:

- Eliminate

- Substitute

- Engineering Controls

- Management Controls (Administrative / Behavioural Controls)

- Personal Protective Equipment

5.4	Laser Controls – Optical Hazards

The simplest rule to follow to avoid eye injury is not to look directly into a laser beam or its specular reflection, regardless of the laser’s power or classification or the laser eyewear being worn. A Maximum Permissible Exposure (MPE) should be calculated for laser sources present in a laser system based on:

- The radiated wavelengths

- Output powers or energies

- Pulse duration and pulse repetition rate (if appropriate).

Engineering and administrative controls should be used to keep exposures below the MPE whenever practicable. Skin protection and laser eyewear should be used only where engineering and administrative controls are impractical.

5.4.1	Laser Alignment

Approximately 60% of laser accidents in research settings occur during the alignment process. Laser alignment guidelines to help prevent accidents include:

- Restricted access, unauthorised personnel must be excluded from the room or area.

- The wearing of laser protective eyewear when appropriate.

- The training and instruction of Class 3B/4 laser users.

- Instruction to remove watches and reflective jewellery before any alignment activities begin.

- The lowest possible / practical power must be used during alignments.

- The use of a He-Ne or CW diode alignment laser, when possible, for a preliminary alignment.

- Identifying individual responsibilities – the individual who moves or places an optical component on an optical table is responsible for identifying and terminating each and every stray beam coming from that component.

- Identifying when the beam is directed out of the horizontal plane.

- Checks on the stability and rigidity of all optical mounts, beam blocks and stray beam shields.

- Use of beam paths at a safe height, below eye level when standing or sitting and not at a level that tempts one to bend down and look at the beam.

5.5	Control of Non-optical Hazards

Many hazards (other than from laser radiation) that can be found in the laser are must be adequately assessed and the risks controlled. The manufacturer’s safety guidance material should help in identifying most of the associated hazards. The main non-optical hazards include:

- Electrical – high voltages and capacitors used with pulsed lasers can present a serious hazard particularly during servicing;

- Collateral radiation – this could include x-rays, UV, RF, visible and IR radiation;

- Fume – can be released from the action of high power lasers used in materials processing and surgery;

- Hazardous substances – substances used in dye and excimer lasers can be toxic and carcinogenic, cleaning solutions may also be hazardous;

- Cryogenic liquids – used with high-powered lasers can present a burning hazard, possible oxygen depletion hazard and possibly an explosion hazard from over-pressure of gases on a closed system;

- Fire and Explosion – high powered (Class 4) lasers can ignite materials and even relatively low-powered lasers (> 35mW) can cause explosions in combustible gases and dusts;

- Mechanical hazards – from gas cylinders, trailing cables and water hoses, cuts from sharp objects, handling difficulties with large work pieces;

- Noise – from discharging capacitor blanks, from some pulsed lasers and from some air-cooled lasers.

Other hazards may also arise from the environment in which the laser is used – adverse temperature and humidity, low light-level conditions, mechanical shock and vibration, interruptions to the power supply, computer software problems and ergonomic problems caused by poor design of the layout of equipment. Issues such as cleaners or maintenance staff inadvertently disturbing equipment or unsupervised access must also be addressed.

5.6	Management Controls

Steps should be taken to change behaviour so that people do not inadvertently or deliberately put themselves at risk. Examples include:

- Training

- Designating Laser Controlled Areas where exposure can exceed the ELV for the particular wavelength of laser radiation

- Signage and warning lights

- Preventing unauthorised access to a controlled area

- Controlling access to keys

- Keeping beam paths as short as possible

- Containing the beam in fibres, flight tubes or other forms of containment

- Written Safe Operating Procedures for high risk activities including maintenance and alignment of the beam

- Controlling the use of equipment or the wearing of items that may give rise to specular reflections

- Local rules

- Use of alignment aids during routine maintenance to realign the beam path e.g. use of a lower power sighting laser or use of mask or target.

- Displaying Emergency Action notices as laminated sheets near to the entrance to any Designated Laser Controlled Area to give instruction on how to obtain emergency assistance and medical help.

5.7	Laser Accident

Some common unsafe practices that are causes of preventable laser accidents are:

- Lack of pre-planning and failure to follow safety protocols.

- Misaligned optics and upwardly directed beams – pay particular attention to periscopes and reflections from windows and beam splitters / combiners.

- Available eye protection not used particularly during alignment procedures.

- Wearing the wrong eyewear.

- Bypassing of door interlocks and laser housing interlocks.

- Insertion of reflective materials into beam paths.

- Lack of protection from non-beam hazards.

- Improper methods of handling high voltage.

- Operating unfamiliar equipment.

5.8	Laser Incident Grab Sheet

In the event of an ocular injury occurring from exposure to laser radiation, it is vital that an eye examination is carried out as quickly as possible by an ophthalmologist at the nearest casualty eye clinic.  It is also essential to provide the ophthalmologist with as much information about the laser radiation as possible.  To assist in these matters, every laser system of Class > 2 should have a completed grab sheet posted near the system (example in Appendix 9) In the event of an incident this sheet can be “grabbed” and presented to the ophthalmologist at the time of the eye examination.

5.9	Assessing residual risk and recording the results

In most circumstances, after introducing control measures, one should be able to assess the residual risk as being low. One then needs to produce written Systems of Work and make it available to all users so that they are aware of all protective measures they should be taking and the procedures they should be following. (See example risk assessment in Appendix 5).

It should be noted that with the changing nature of experimental work it is important that the risk assessment and operating procedures are routinely reviewed and, most importantly, reviewed prior to any significant change.

- Personal Protection

Whenever there is a risk of laser exposure to levels above the specified MPEs, personal eye-protection is one of the commonest and important elements of personal laser protection. Personal eye-protection should be regarded as a last line of defence against exposure to laser radiation; it can be adopted only after a full safety evaluation has been carried out and other means of affording protection have been considered and rejected.

Its use should not be regarded as a convenient alternative to proper engineering controls or thorough hazard assessments.

When choosing personal eye-protection, it is important to consider not only the ability of the eyewear to attenuate the incident radiation but also to have a damage threshold high enough to withstand the maximum possible exposure long enough for avoiding action to be taken, and to take into account comfort and visual light transmission (mean percentage of visible spectrum that is not filtered by the eyewear). Special consideration needs to be made when selecting eyewear for femtosecond and picosecond pulsed lasers.

Laser protective eyewear should never be relied on to provide protection against deliberate exposure to a laser beam but should be regarded as a means of providing some protection against accidental exposure.

Eye protection which is designed to provide adequate protection against specific laser wavelengths should be used in all hazard areas where Class 3R laser products emitting energy outside of the 400nm – 700nm wavelength range, Class 3B or Class 4 lasers are in use. Exceptions to this are:

- When engineering and administrative controls are such as to eliminate potential exposure in excess of the applicable MPE;

- When, due to the unusual operating requirements, the use of eye protection is not practicable. Such operating procedures should only be undertaken with the approval of the Departmental Laser Safety Officer.

When working with Class 4 lasers (and some Class 3B devices emitting in the UV) skin protection may also be required.

Special attention has to be given to the resistance and stability against laser radiation when choosing eyewear or protective clothing for protection against Class 4 lasers.

Personal protective equipment should be personal, i.e. it should be appropriately cleaned between users, or each user has their own.

- Medical Supervision, Emergency Eye Examinations and Accidental Exposures

Eye examinations for laser users are not recommended as part of a safety programme. The value of routine examinations for Class 3B/4 laser users has been reviewed and it is generally accepted that routine examinations are of little value.

It is more important to have a procedure in place if there has been an apparent or suspected ocular exposure. A medical examination by a qualified specialist needs to be carried out as soon as possible. In the event of an accident or incident involving suspected injury to the eye(s), an emergency examination should be carried out as soon as possible and within 24 hours.

The most appropriate Accident and Emergency Department (RVH) which deals with eye injuries has been identified and the injured party should be taken here. Suitable arrangements should be in place to ensure that all persons working with Class 3B/4 lasers are aware of the action to take in the event of an accident / incident.

Each Class 3B/4 laser should have a card or proforma that can be taken with the casualty to hospital. An example of such a card and the information that will be required in the event of an accident / incident is given in Appendix 9.

In the event of an eye injury caused by an individual staring down the beam of a powered laser, the emergency arrangements for Class 3B/4 lasers should be followed.

Where an emergency eye examination is required, the Laser Safety Officer and Laser Safety Supervisor will carry out detailed investigation of the accident / incident.

In the event of a skin injury, i.e. thermal burn, this can be treated as would any other burn.

All accidents and incidents, whether involving an emergency examination or not, must be reported promptly to the Occupational Health and Safety Service using the appropriate Accident Report Form (AC1).

Depending on whether an injury has been sustained, there may be a requirement to notify the Health and Safety Executive for Northern Ireland (HSENI) under the Reporting of Injuries, Diseases and Dangerous Occurrences Regulations (RIDDOR).

There is no statutory requirements under the Control of Artificial Optical Radiation at Work Regulations (NI) 2010 for medical surveillance for the eye. However, appropriate medical supervision is required of someone is exposed to laser radiation in excess of the MPE (termed Exposure Limit Value in the Regulations).

- Laser Pointers

Small lasers are readily available and some are used for presentation purposes as pointers. These laser pointers are normally classified at a level above Class 1 and therefore in some circumstances can cause harm, particularly by staring into the beam. In the past they were only available in red wavelengths and had output powers up to and sometimes above 5mW. Nowadays, devices are available emitting green wavelengths where the eye is more sensitive and lower powers are all that is needed so that laser pointers should now be Class 2 devices (output < 1mW). Normally the eye’s natural aversion response, including the blink reflex, affords protection to short duration accidental exposure. Where laser pointers are used, instructions on their safe use should be readily available. Appendix 8 is an example of the instruction that should be provided.

| Appendix 1 – Laser Inventory School / Centre: …………………………………………………….. Laser Safety Supervisor: ………………………………… Date: ……………………………………………… Manufacturer Laser Type Class Wavelength Max Output (W or J) CW / Pulsed Pulse Length Model No. Serial No. Location |  |  |  |  |
| --- | --- | --- | --- | --- |
| Example Uniphase He-Ne 3B 632.8nm 5mW CW - 1105P 372422 EC 21 |  |  |  |  |

Appendix 2: Registration form for Laser Users

REGISTRATION OF USERS OF LASER EQUIPMENT OF CLASS > 2

Before commencing work with ionising radiations it is necessary to register with the University’s Laser Safety Officer (LSO).

In addition, it is vital that you are familiar with your Departmental Local Rules concerning work with laser equipment.  If you have any doubts or queries please contact either your Departmental Laser Safety Supervisor (DLSS) or the University’s LSO at Ext 4612.

Please complete the following details and forward the form to your Departmental LSS who will send it to the LSO under sealed envelope.

Surname: __________________________	Forenames: ________________________________

Sex: ______________________________	Date of Birth: _______________________________

School/Department: ________________________________________________________________

Laboratory/Room Number: ___________________________________________________________

Details of Laser System to be used: ____________________________________________________

_________________________________________________________________________________

_________________________________________________________________________________

Please read the following carefully and sign:

I have read, and am fully familiar with and will carefully follow the Departmental Rules for work with laser systems.

I wish to be registered as a Laser Worker and am confident that my practical ability will allow me to work safely and without danger to myself or my colleagues.

Signature: ________________________	Project Supervisor’s Signature: _________________

Date: ____________________________	Date:  ____________________________________

Appendix 3

Use of Class 1M, 2M and 3R Lasers

Risk Assessment

Without the use of magnifying optics, 1M devices do not pose an eye hazard, neither do 2M or Class 2 devices as long as you do not stare into the beam (eye protection is normally afforded by the aversion responses). An eye hazard is possible if there is:

- Exposure in excess of more than 0.25 seconds from Class 2/2M lasers

- Exposure to Class 1M/2M

- Of if Class 3R lasers are viewed directly.

The risk of eye injury is low and there is no skin or fire hazard.

1. Location of the Work Activity

| School / Centre |  |
| --- | --- |
| Location of work (building and room number) |  |
| Principal Investigator / Researcher /Supervisor |  |
| Brief description of work activity |  |
| Assessor |  |
| Date of Assessment |  |

2. Laser Specifications

| Model: |  |
| --- | --- |
| Maximum Power: |  |
| Wavelength Range: |  |

3. Non Optical Hazards

Detail the significant risks and the control measures necessary for any non-optical hazard identified and any optical hazard from 1M/2M devices.

| Significant Hazards | People at risk | Existing Controls | Risk Rating | Additional Controls required | Remaining level of risk |
| --- | --- | --- | --- | --- | --- |

4. Control Measures

Avoid eye level and do not expose users or others to the beam.

N.B. Modified 1M/2M devices may need to be reassessed as a higher classified laser

- Follow manufacturer’s safety instructions

- Take care when operating the laser system

- Keep the laser ‘on’ only when necessary

- Restrict unauthorised use

- Terminate the beam at the end of its useful path

- Do not point at or towards persons deliberately

- Do not point at mirrored surfaces that may cause unplanned reflections

- Never look into the laser aperture

- Never look directly or stare into the beam/beam aperture when on

- Never allow unauthorised use

- Do not use optical viewing aids.

5. Contact

Name:							Phone Number:

Appendix 4

Use of Class 3B and 4 Lasers

Risk Assessment

Class 3B and Class 4 lasers are capable of causing eye injury to anyone who looks directly into the beam or its specular reflections. In addition, diffuse reflections of a high-power (Class 4) laser beam can produce permanent eye damage. High-power laser beams (Class 4) can burn exposed skin, ignite flammable materials and heat materials that release hazardous fumes, gases, debris or radiation. Equipment and optical apparatus required to produce and control laser energy may also introduce additional hazards associated with high voltage, high pressure, cryogenics, noise and other forms of radiation, flammable materials and toxic fluids. Thus, each proposed experiment or operation involving a laser must be evaluated to determine the hazards involved and the appropriate safety measures and controls required.

1. Location of the Work Activity

| School / Centre |  |
| --- | --- |
| Location of work (building and room number) |  |
| Principal Investigator / Researcher /Supervisor |  |
| Brief description of work activity |  |
| Assessor |  |
| Date of Assessment |  |

2. Laser Identification and Specifications

Complete the following table; list all lasers including low power alignment lasers

|  | Laser 1 | Laser 2 | Laser 3 |
| --- | --- | --- | --- |
| Type: |  |  |  |
| Manufacturer: |  |  |  |
| Model: |  |  |  |
| Serial Number: |  |  |  |
| Maximum Power: |  |  |  |
| Maximum Pulse Energy: |  |  |  |
| Wavelength Range: |  |  |  |
| Wavelength Used: |  |  |  |
| Power Used: |  |  |  |
| Pulse Energy Used: |  |  |  |
| Pulse Length: |  |  |  |
| Pulse Repetition Rate: |  |  |  |
| Beam Diameters (x,y) |  |  |  |
| Beam Shape (e.g. ellipse, box, circular) |  |  |  |
| Beam Divergence (x,y) |  |  |  |
| LASER CLASSIFICATION |  |  |  |

3. Description of activity or research project

Provide a brief description of the laser set-up, its purpose and operational parameters.

| Duration of Activity / Project Is the work ongoing or for a limited period of time? |  |
| --- | --- |

4. Identification of hazards additional to the laser hazard(s)

Electrical Hazards Most lasers contain high-voltage power supplies and often large capacitors / capacitor banks that store lethal amounts of electrical energy. Are any special precautions/procedures required? Yes / No

Laser Dyes Laser Dyes are often toxic and/or carcinogenic chemicals dissolved in flammable solvents. Are laser dyes used? Yes / No If yes, give details:

Compressed and Toxic Gases Hazardous gases may be used in laser applications, i.e., excimer lasers (fluorine, hydrogen chloride) Are compressed gases and/or toxic gases used? Yes / No If yes, give details:

Cryogenic Fluids Cryogenic fluids can create hazardous situations. Adequate ventilation must be provided. Are cryogenic fluids used? Yes / No If yes, give details:

Fumes/Vapours/Laser generated air contaminants from beam/target interaction When laser beams are sufficiently energised to heat up a target, the target may vaporise, creating hazardous fumes or vapours that may need to be captured or exhausted. Is there a potential for fumes/vapours/laser generated air contaminants? Yes / No If yes, give details:

UV and Visible Radiation/Plasma Emissions UV and visible radiation may be generated by laser discharge tubes, pump lamps or plasmas. The levels produced may be an eye and skin hazard. Is there potential for significant UV/visible radiation? Yes / No

Explosion Hazards High-pressure arc lamps, filament lamps and capacitors may explode if they fail during operation. Laser targets and some optical components also may shatter if heat cannot be dissipated quickly enough. Is there an explosion hazard? Yes / No

Ionising Radiation (X-rays) X-rays can be produced from two main sources, high voltage vacuum tubes of laser power supplies such as rectifiers, thyratrons and electric discharge lasers. Any power supplies that require more than 15 kV may produce x-rays. Is there an ionising radiation hazard? Yes / No

Other hazards not identified above Please specify

5. Persons who may be at risk

6. Measures to reduce level of risk (laser hazard)

Are open or partially enclosed beams used during the following: 1. Initial setting up and beam alignment Yes / No / N/A 2. Addition of new optical elements/lasers Yes / No / N/A 3. Day to day operations Yes / No / N/A 4. Servicing Yes / No / N/A

Are there protocol/procedures to control risks from the ocular (and if applicable, skin) hazard during the following? 1. Initial setting up and beam alignment Yes / No / N/A 2. Addition of new optical elements/lasers Yes / No / N/A 3. Day to day operations Yes / No / N/A 4. Servicing Yes / No / N/A

List the operating protocols with references/dates/location: (All open beam work must have an appropriate protocol / Operating Procedure)

7. Instruction / Training

Authorised laser users must receive appropriate training and instruction

Specify the instruction and training arrangements

A list of authorised laser users is to be displayed

8. Personal Eye Protection

Detail how scale number was determined

| Number Available | Location | Manufacturer | Scale Number | Wavelength |
| --- | --- | --- | --- | --- |

9. Assessment of Risk (Associated Hazards identified in Section 4)

Detail the significant risks and the control measures necessary (i.e. by reference to protocols/procedures or safety manual). For hazardous substances, specify the location of the appropriate COSHH assessments.

| Significant Hazards | People at risk | Existing Controls | Risk Rating | Additional Controls required | Remaining level of risk |
| --- | --- | --- | --- | --- | --- |

10. Monitoring

Supervisors should monitor that users are complying with procedures as should the School/Centre and Departmental Laser Safety Supervisors by carrying out periodic checks.

Performance of control measures:

11. Review

Where new lasers or components are introduced then these changes need to be assessed; protocols may need to be modified. A review would also be required where a new laser worker starts; ensure that they are informed of the relevant risks and protocols.

Enter the date or circumstances for review of assessment (3 years or the length of the particular project/worker if shorter).

12. Emergency Action

To Control Hazards To stabilise situation e.g. turn off power source, etc.

To Protect Personnel Evacuation, protection for personnel, Special first aid

To Render Site of Emergency Safe Power, Ventilation

13. Emergency Contact

Name:							Phone Number:

Appendix 5

Scheme of Work with Procedure (Example 1)

School / Centre	…………………………………………………………………………………

Laboratory / Room / Activity	………………………………………………………………..

These guidelines have been prepared by	……………………………………………………….

Before commencing work with Class 3B/4 lasers, you must read this document, and sign the sheet at the end to confirm that you have understood the content and that you agree to abide by the protocols contained herein.

Purpose and Structure of this document

The principal aim of this document is to outline the elements of good laser practice as they apply specifically to experiments currently being undertaken in the above laboratory. General aspects of laser safety are covered in sections of the manuals accompanying the lasers.

The document is structured as follows: at the top level (this sheet) an overall description is made of laser research activity in this laboratory. The user is then referred to a number of accompanying documents under two headings: Laser types in use and Safety Protocols.

Description of Activity

Description of the activity / experiment and its purpose

Four types of situation have been identified which require separate safety protocols, where appropriate:

- Setting up;

- Adding new elements to;

- Day-to-day operation of the experiment; and

- Maintenance of the laser system.

Separate sections dealing with each of these are to be included where appropriate.

Lasers in Use

Reference to completed Hazard / Risk Assessment

As a general rule all Class 3B/4 laser emissions are capable of causing severe eye damage if viewed directly, or as a specular (i.e. mirror-like) reflection. Control measures (careful planning, beam pipes, blocking of reflections, safety eyewear) must be taken to avoid this.

Authorised users of the above lasers are (to be named):

Supervisor(s):

Student(s):

Other(s):

There should be reference to this guidance in particular the General Overview on Laser Hazard / Risk Assessment and Maximum Permissible Exposure Calculations.

Laser Safety Protocol: Setting up

Definition

Setting up applies to the initial installation of a new experiment and to major changes such as the addition of a new type of laser system, or, for example, a complete change of beam paths.

Protocol / Scheme of Work

Planning:

- The installation or changes should be discussed with a supervisor prior to operation of Class 3B/4 laser systems.

- In the case of a completely new experiment, the University Laser Safety Officer must be consulted and invited to visit the lab.

- The laser beam paths and associated optics should be planned to minimise the possibility of stray reflections.

- Termination of each main laser beam should be planned.

- Consideration should be made on how alignment is to be carried out. Where practicable, the use of cameras for viewing and remote adjustment should be promoted.

- Provision of suitable laser safety eyewear should be addressed at this stage.

Initial safety checks:

- Before starting the Class 3B/4 lasers, beam paths should be inspected for any objects that should not be there.

- Laser warning signs should be activated, where installed and unauthorised persons excluded and doors closed.

- Alignment may be carried out by one or at most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Appropriate personal eye-protection should be worn where required.

Initial alignment and suppression of stray reflections:

- Initial laser beam alignment should be performed with a Class 1 or 2 alignment laser (e.g., He-Ne or small cw diode laser). Remember that the final beam path may differ slightly due to dispersion (i.e., the beam path may be slightly wavelength-dependent).

- At this stage, each and every optic element in the beam path must be analysed for stray reflections. Initially this can be done by predicting the likely path of specular (i.e., non-diffuse) reflections and the actual reflections of the Class ½ alignment laser may also be used to help identify stray reflections.

- Suitable beam blocks, opaque at the appropriate wavelengths, must then be installed to block all these stray reflections.

- ‘Beam pipes’ should be installed at this stage to cover longer runs of laser beam, and especially any beams that leave the confines of the laser table. It is recognised that there may be some places where beam pipes are inappropriate, e.g. when the distance between optics is very short. Beam pipes should be designed to allow limited access to the beam for alignment checking without removal.

Alignment using Class 3B/4 lasers at low power

- The next stage of alignment using the Class 3B/4 lasers may be carried out only after obtaining the verbal permission of a supervisor.

- Alignment may be carried out by one or at most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Under no circumstances must direct viewing of the laser beam be attempted even if the beam has been attenuated.

- All optics should be checked for damage, and the stability of optics mounts verified prior to operation of laser.

- This next stage in alignment should be carried out using the lowest possible laser energy (e.g. operating a Nd:YAG laser on fixed-Q) at which it is possible to visualise the laser beam in an appropriate fashion. The method of visualisation is dependent on the wavelength: for UV or visible light, the beam can be viewed on a fluorescent card. An invisible infrared beam may be visualised using LCD heat sensitive paper or possibly using burn paper or a laser power meter.

- In the case of UV or IR beams, appropriate laser safety eyewear should be worn during the alignment procedure at all times when the laser pulse energy exceeds the MPE: note that it should not block the wavelength-shifted visible fluorescence (UV) or the heat effect on LCD paper or burn paper (IR), which can then be used to visualise the beam.

- Visible or multi-wavelength alignment may have to be carried out without laser safety eyewear, as it would otherwise be impossible to visualise the laser beam on a card. In this case, extra caution must be exercised by the operator(s) after appropriate consideration of alternatives and assessments of risks. The laser(s) should be operated below the MPE if possible and in any case at the lowest practicable pulse energy. Blocking of possible stray reflections must be double-checked prior to carrying out this stage.

- Alignment of each laser beam to variable diameter apertures (iris diaphragms) should be employed where possible to minimise the necessity for multi-wavelength alignment.

- Further alignment at full power may be carried out in accordance with the protocol outlined under ‘day-to-day operation’.

Laser safety protocol: Adding new elements

Definition

Adding new elements applies to the introduction of any new optic into the beam path of a Class 4 laser such as a lens or filter.

Protocol / Scheme of work

Planning:

- The placement of additional optics should be planned to minimise the possibility of stray reflections.

- Beam blocks should be devised to terminate any unavoidable stray reflections.

Initial safety checks:

- Before starting the Class 3B/4 lasers, beam paths should be inspected for any objects that should not be there and beam pipes should be replaced if necessary.

- Laser warning signs should be activated, unauthorised persons excluded and laboratory doors closed.

- Alignment may be carried out by one or at most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Appropriate laser safety eyewear should be worn if practicable. Visible or multi-wavelength alignment may have to be carried out without laser safety eyewear, as it may otherwise be impossible to visualise the laser beam on a card. In this case extra caution must be exercised by the operator(s).

- All optics should be checked for damage, and stability of optics mounts verified.

Initial alignment and suppression of stray reflections:

- Once a new optic is in place, initial alignment should be performed with a Class 1 or 2 alignment laser (e.g., He-Ne or small cw diode laser). For simple optics it may be judged sufficient to proceed to the next step without using a Class 1 / 2 alignment laser.

- The new optic element in the beam path must be analysed for stray reflections. This can be done by predicting the likely path of specular (i.e., non-diffuse) reflections. The actual reflections of the Class 1/2 alignment laser may also be used to help identify stray reflections.

- Suitable beam blocks, opaque at the appropriate wavelengths, must be installed to block all these stray reflections.

- Any effect ‘downstream’ of the new optic should be checked. ‘Beam pipes’ should be re-installed at this stage.

Alignment using Class 3B/4 lasers at low power:

- This may now be carried out in accordance with the procedure outlined under ‘setting up’ with the exception that explicit permission of a supervisor is not deemed necessary for addition of a simple optical element. (Anything more complex should be taken as ‘setting up’ and the protocol followed accordingly).

Laser safety protocol: Day-to-day running

Definition

Day-to-day running applies to the operation of Class 3B/4 lasers under all circumstances except setting up or addition of a new optic element. It includes initial, minor realignment of laser beams at the beginning of an experimental run and ‘tweaking’ of alignments during an actual experiment.

Protocol / Scheme of work

Initial safety checks:

- Before starting the Class 3B/4 lasers, beam paths should be inspected for any objects that should not be there, and beam pipes should be replaced if necessary.

- Laser warning signs should be activated, unauthorised persons excluded, and laboratory doors closed.

- Alignment may be carried out by one or at the most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Appropriate laser safety eyewear should be worn if practicable. Visible or multi-wavelength alignment may have to be carried out without laser safety eyewear, as it may otherwise be impossible to visualise the laser beam on a card. In this case extra caution must be exercised by the operator(s).

- All optics should be checked for damage, and stability of optics mounts verified.

Check using Class 3B/4 lasers at low power:

- Before turning on full power, the beam path of each laser should be verified in turn, using the lowest possible pulse energy and visualising the beam in an appropriate fashion (e.g., on fluorescent card).

- Under no circumstances must direct viewing of the laser beam be attempted even if the beam has been attenuated. There must be no exceptions to this rule.

Minor realignment (‘tweaking’) with lasers running at full power:

- During an experimental run, it will sometimes be necessary to re-optimise the alignment to recover lost signal. Of necessity, this can only be carried out at full power, with all lasers on. Extra caution should therefore be exercised.

- All beam guards/pipes and blocks for stray reflections should remain in place during this procedure. Beam pipes should be designed to allow limited access to the beam for alignment checking without removal.

- It is especially important to wear appropriate laser safety eyewear when visualising laser beams at full power. As before, in a visible or multi-wavelength experiment this may not be practicable, and extra caution should therefore be exercised.

- It may be possible (and indeed, preferable) to apply minor ‘tweaks’ to the alignment using the experimental signal as a guide. In this case it is not necessary to visualise the laser beams.

- Cameras for remote viewing and the incorporation of remote adjustment aids should be promoted and used where reasonably practicable.

Laser safety protocol: Servicing

Definition

Servicing is the performance of those procedures or adjustments described in the manufacturer’s service instructions which may affect any aspect of the product’s performance. It can include activities such as the removal and reinstallation of optics for cleaning, the changing of laser dyes, the changing of flash lamps, and the installation of new optics inside the laser cavity. Entry into the laser enclosure potentially exposes the laser worker to additional non-optical hazards, for instance those associated with high voltages and toxic chemicals, in addition to accessing high energy laser beams which are normally enclosed. The activities are very diverse, and specific protocols will need to be prepared for each laser, this document provides some general guidelines on the planning of Maintenance.

Protocol / Scheme of Work

Planning

- Before commencing the work, the manual for the laser system should be consulted to identify the recommended procedure.

- In the case of anything other than routine maintenance, and/or when the laser manual does not give a procedure, the advice of a laser technician should be sought. Some procedures should only be conducted by an experienced laser technician.

- The hazards associated with the procedure should be assessed, the control measures reviewed, and the conclusions recorded. In the case of some regular maintenance procedures (such as changing the dye solution in a dye laser), reference to an existing protocol may well suffice.

- Work involving the alignment of a laser beam inside a laser enclosure, for instance introducing the pump laser into a dye laser, can lead to an increased laser radiation exposure risk, since part of the beam path of a normally enclosed, and potentially very high power beam is likely to be open. The protocol for Setting Up should be consulted.

A signature sheet will act as a list of authorised users and a record that persons have been given information and agree to follow the procedures laid out.

Signature Sheet

This sheet must be signed by all workers before commencing work with Class 3B/4 lasers. By signing below, you confirm that you have read this document, you understand the requirements and that you agree to abide by the protocols and guidelines contained herein.

| Name (in block capitals) | Signature | Date |
| --- | --- | --- |

Simple Set up (Example 2)

Laser Scheme of Work			Class 4, 3 3, 635nm Diode Laser (CW)

In Location, Department, Premises

Experiments:					Photodynamic therapy in vitro and vivo.

Person responsible:				Prof A N Other

Authorised operatives:			Prof A N Other

Dr L Aser

General Precautions

- Only the above names personnel are permitted to use this laser. They must be familiar with the manufacturer’s and the University’s safety information (including this scheme of work).

- When not in se the key to the laser (or other means of controlling access) should be kept separately from the laser, to prevent unauthorised use.

- All beam pats shall be kept as short as possible and enclosed whenever reasonably practicable.

- The area in which this laser is used should be a designated laser area, and have the appropriate warning notice on the door.

Specific considerations relating to this experiment

The laser should be sited at least 1.1m away from the window (NOHD is 1.09m) and the blinds should be closed.

Before switching the laser on:

- The fibre optic head must be securely mounted and pointing vertically down on to a non-reflective surface.

- The appropriate safety eyewear should be worn that gives at least OD 2.3 (L3) at 635nm. Only those wearing safety eyewear are permitted in the operating area.

- The ‘Laser in Use – Do Not Enter’ sign should be posted on the entrance door.

The laser must be aligned using the lowest practicable power setting.

The laser should not be left unattended unless a remote interlock connector linked to the door prevents unauthorised access.

Because of the diverging nature of the beam the eye and skin hazard from these laser experiments is low. The only other hazard identified with this work is from the cabling. All cabling should be tied/taped back away from the operating are to prevent any tripping hazard and damage to the alignment of the fibre optic.

Copies of this Scheme of Work should be issued to all personnel using the laser, and should also be posted in the laboratory.

Departmental Laser Safety Supervisor………………………….	Date………………………

Appendix 6

Summary of Warning and Protective Control Measures

| Class | Protective Control Measures |
| --- | --- |
| 1 | No protective control measures for normal use (NB special precautions may be needed for service work on embedded laser products). |
| 1M | Prevent direct viewing with magnifying optics. (NB fitting external optics that decrease beam divergence may affect classification).* |
| 2 | Do not stare into beam. Do not direct the beam at other people or into public areas. |
| 2M | Do not stare into beam. Do not direct the beam at other people or into public areas. Terminate beam at end of useful path with non-specular beam stop. Prevent direct viewing with magnifying optics. (NB fitting external optics that decrease beam divergence may affect classification).* |
| 3R | Prevent direct eye exposure to the beam. Do not direct the beam at other people or into public areas* |
| 3B and 4 | Class 3B and Class 4 laser products should not be used without first carrying out a risk assessment to determine the protective control measures necessary to ensure safe operation. Where reasonably practicable engineering means should be used to reduce the laser class to a totally enclosed Class 1 laser product. The use of any Class 3B or Class 4 laser without an interlocked enclosure will require a written Scheme of Work. Even with an enclosure, written procedures will be necessary if the user is involved in any alignment procedures that require over-riding of interlocks. Class 3B and Class 4 laser products required the control of access to the area where the laser is operated by the use of a remote interlock, the use of key control, emission indicators, beam shutters, removal of reflecting surfaces that could be struck by an errant beam, beam enclosures wherever practical, the use of eye protection and protective clothing as appropriate, training of staff and the appointment of a Laser Safety Officer. |

*Classes 1M, 2M and 3R may also require training of staff, care with beam paths and specular reflections – see BS EN 60825-1 and PD IEC/TR 60825-14 for more details.

Special attention should also be given to other non-optical hazards such as risk of electric shock, hazardous chemicals, cryogenic liquids and flying debris from targets to name but a few. It is often the non-optical hazards that pose the greatest risk – one could be blinded in one eye from a powerful laser but electrocution could be fatal. Some non-optical hazards may be present even with Class 1 laser products.

Appendix 7

Summary of BS EN 60825-1

Manufacturers and User Requirements

| 1 | Remote Interlock | Connection provided by the manufacturer for door or enclosure interlock for Class 3B and Class 4 lasers. |
| --- | --- | --- |
| 2 | Safety Interlocks | Require for access panels on Class 3R, 3B and 4 laser systems. |
| 3 | Key Controls | A key or similar device is required to control unauthorised operation of Class 3B and Class 4 lasers. |
| 4 | Emission Indicator | An audible or visible indicator should be provided by the manufacturer for each Class 3R laser (except wavelengths 400 – 700nm) and each Class 3B and Class 4 laser system. |
| 5 | Beam stop or attenuator / shutter | Should be provided by the manufacturer for each Class 3B or Class 4 laser system. |
| 6 | Beam termination | The user should ensure that all beam paths are terminated at the end of their useful path. (Does not apply to Class 1 devices). |
| 7 | Beam level | Avoid eye level. |
| 8 | Beam enclosure | To guard against specular reflections from Class 3R, Class 3B and Class 4 lasers – can mean anything from screening the experimental area or piping the beam up to a total enclosure. |
| 9 | Eye protection | Required for all open beam work with invisible beam Class 3R and all Class 3B and Class 4 Devices. |
| 10 | Protective clothing | Mainly required for Class 4 lasers but be careful with Class 3B UV lasers as well, may need fire resistant material for some lasers. |
| 11 | Eye examinations | Only required after an accident but mat be important to people with poor eyesight working with Class 3B or Class 4 lasers. |
| 12 | Training | Required for people working with any Class 3 or Class 4 laser and any modified Class 1M or Class 2M devices. |
| 13 | Laser labels | Required for all lasers except low power Class 1 (though need not be directly affixed if the size of the laser product does not permit this). |
| 14 | Door / Area signs | Required for Class 3B and Class 4 lasers indoors and also for Class 1M, 2M and 3R if used outdoors. |

Appendix 8

The Safe Use of Laser Pointers

Small lasers are commonly available and some are used for presentation purposes as pointers. In the past, laser pointers were only available in red wavelengths and had output powers up to and sometimes over 5mW. Nowadays, devices are available emitting green wavelengths, where the eye is more sensitive and lower powers are all that is needed, so that laser pointers now only need to be Class 2 devices (output <1mW).

NB	Class 1 laser products are normally safe.

Class 2 and Class 3R products are not hazardous under certain conditions, however they can cause harm to the eyes, particularly if the beam is stared into.

Class 3B laser pointers must not be used.

Normally the eye’s natural aversion response affords protection to short duration accidental exposure to Class 2. Class 1 or Class 2 laser pointers are the recommended choice where a laser pointer is necessary and are the only type that should now be purchased.

Laser pointers should only be used as a pointing device and securely stored when not in use. Persons who use laser pointers should ensure that they are aware of potential hazards and they should comply with the basic instructions below.

Instructions for Use

When operating laser pointers, users must ensure that they use them in a safe manner and do not expose themselves or others to the beam. Laser pointers are not to be modified in any way.

- Follow the manufacturer’s safety instructions

- Take care when operating the laser pointer

- Keep the ‘on’ button depressed only when necessary

- Do not keep the ‘on’ button depressed when not pointing at the screen

- Do not point at or towards the audience

- Do not point at mirrored surfaces

- Never look into the laser aperture

- Never look directly or stare into the beam/beam aperture when on

- Never allow unauthorised use, especially by children

Further guidance on laser pointers can be found at:

http://www.hpa.org.uk/web/HPAweb&HPAwebStandard/HPAweb_C/1195733794576

Appendix 9

Laser Incident Grab Sheet

In the event of an ocular injury resulting from a laser incident, an examination should be carried out by an ophthalmologist at the nearest Casualty Eye Clinic. The examination should be carried out as soon as possible after the injury and the information contained in this sheet should be made available to anyone treating the injury.

Do not drive yourself; get a friend or colleague to take you.

Out of hours:	Contact Security or Emergency Services

- State Building and School / Centre

- Location and nature of incident / accident

- Request assistance to take the casualty to the RVH

- Complete the information below.

| Emergency Ophthalmic Examination LASER Exposure |
| --- |
| Report to: | Eye Casualty, Ophthalmology, Level 8 Outpatients Centre, Royal Victoria Hospital, Belfast Tel: 02890240503 |
| Laser Details: |
| Type: Wavelength: Power Output (CW), or Pulse Energy, Duration and Rate (pulsed): Laser Classification: | Continuous Wave / Pulsed* …………………………...nm ………………………………. ………………………………. ………………………………. |
| School / Centre: |  |
| Name of Injured Party: |  |
| Exposure Details: |
| Circumstances of accident / injury: Time / Date of Injury: Eye affected: Laser settings at time of injury: Were protective goggles worn? | ……………………………………………………………… ……………………………………………………………… Left / Right / Both* ……………………………………………………………… Yes / No* *delete as appropriate |
| Report Accident / Incident to the University Occupational Health and Safety Service on Ext: 4612/4613 during normal working hours and ensure that an AC1 form is completed. |

Appendix 10

Laser Signs and Labels

Designated Laser Areas

The points of access to areas in which Class 3B or Class 4 laser products are used must be marked with warning signs complying with the Health and Safety (Safety Signs and Signals) Regulations (Northern Ireland) 1996.

The signs shall incorporate the following information:

- Hazard warning symbol

- Highest class of laser in the area

- Responsible person with contact details

Laser Labels

Laser labels are required for all laser products except for low power Class 1 devices. They are designed to give a warning of laser radiation, the class of laser, basic precautions and the laser’s characteristics.

The laser warning uses the same symbol as for the door sign in an appropriate size for the laser to be labelled and should be clearly visible. Supplementary information should be black text on yellow background in accordance with BS EN 60825-1.

Where the size of the laser product does not permit the affixing of a reasonably sized label, a sign should be displayed in close proximity to the laser with all appropriate information on it. Infromation over and above that specified by BS EN 60825-1 is required for Class 1 products that are Class 1 by engineerng design. For these types of laser products it is specified that they are totally enclosed systems and you must give details of the laser enclosed. The BS requirement is just to describe them on the outside as a Class 1 laser product.

Details of working required on explanatory labels are given below:

Class 1 (by engineering design)

No hazard warning label.

Explanatory label bearing the words:		CLASS 1 LASER PRODUCT

A TOTALLY ENCLOSED LASER SYSTEM

CONTAINING A CLASS …… LASER

In addition, each access panel or protective housing shall bear the words:-

CAUTION – CLASS ….. LASER RADIATION WHEN OPEN

With the appropriate class inserted and then followed by the hazard warning associated with that class of laser (see warning statements in following labels).

CLASS 1M

No hazard warning label.

Explanatory label bearing the words:

NB – ‘Optical Instruments’ can be supplemented with either ‘Binoculars or Telescopes’ (for a large diameter collimated beam) or ‘Magnifiers’ (for a highly diverging beam).

CLASS 2

Label with hazard warning symbol.

Explanatory label bearing the words:

CLASS 2M

Label with hazard warning symbol.

Explanatory label bearing the words:

NB – ‘Optical Instruments’ can be supplemented with either ‘Binoculars or Telescopes’ (for a large diameter collimated beam) or ‘Magnifiers’ (for a highly diverging beam).

CLASS 3R

Label with hazard warning symbol.

Explanatory label bearing the words:

For λ 400nm – 1400nm ONLY.

NB – For other λ replace ‘AVOID DIRECT EYE EXPOSURE’ with ‘AVOID EXPOSURE TO BEAM’

CLASS 3B

Label with hazard warning symbol.

Explanatory label bearing the words:

CLASS 4

Label with hazard warning symbol.

Explanatory label bearing the words:

Aperture Labels for Class 3R, Class 3B & Class 4 lasers

Each Class 3R, Class 3B and Class 4 laser product shall display a label close to where the beam is emitted bearing the words ‘LASER APERTURE’ or ‘AVOID EXPOSURE – LASER RADIATION IS EMITTED FROM THIS APERTURE’. This label can take the form of an arrow if this displays more meaning:

Radiation Output and Standards Information

All laser products, except for low power Class 1 devices, shall be described on an explanatory label with details of:

- Maximum output

- Emitted wavelength

- Whether laser beam is visible, invisible or both

- Pulse duration (if appropriate)

- Name and publication date of classification standard

If may be found useful to also put on the labels, details of the type of laser and the lasing medium, although this is not a BS requirement.

Information put on explanatory labels may be combined.

Appendix 11

Optical Hazards and Biological Effects of Laser Radiation

Penetration of laser radiation into the eye

NB Short pulsed high peak-power laser beams are particularly hazardous to the eye, especially at wavelengths that reach the retina, as they deliver a lot of energy in a short period of time that can cause irreversible damage. Near infra-red lasers are also particularly hazardous because you can’t see the beam but it could be focused on the retina and you would only be aware of it after damage had been caused.

Penetration of laser radiation into the skin

In general, the skin can tolerate a great deal more exposure than the eye and less research has been done on damage mechanisms. All wavelengths of laser output with sufficient power density can cause surface burns of the skin and with high-powered Class 4 lasers there could be no warning of this occurring, Near infra-red lasers are again of particular concern because they are more penetrating and can reach the subcutaneous layer and UV lasers are also of concern because of the long-term risk of developing skin cancer.

Appendix 12

Training Record Form – Authorised Laser User

Researcher		………………………………………………………….

Room Number		………………………………………………………….

Laser Used		………………………………………………………….

………………………………………………………….

………………………………………………………….

………………………………………………………….

………………………………………………………….

Attended University training seminar

Registered as a University Laser User

Read University Policy / Guidance for laser safety

Familiar with all hazards within the laboratory and the laboratory’s risk assessment

Trained in use of lasers named above

Familiar with laser Scheme of Work

Signed:

Researcher	………………………………………	Date	……………………………………...

Supervisor	………………………………………	Date	……………………………………...

Return this form to your Department Laser Safety Supervisor

# Text boxes

For the area signs the specifications are quite simple – 50% of the area should be yellow and the width of the black border is 0.06 x the length of the side. A more detailed specification is given for the symbol used in labels, see spec in BS EN 60825-1

LASER RADIATION

DO NOT VIEW DIRECTLY

WITH OPTICAL INSTRUMENTS

CLASS 1M LASER PRODUCT

LASER RADIATION

DO NOT STARE INTO BEAM

CLASS 2 LASER PRODUCT

LASER RADIATION

DO NOT STARE INTO BEAM OR VIEW

DIRECTLY WITH OPTICAL INSTRUMENTS

CLASS 2M LASER PRODUCT

LASER RADIATION

AVOID DIRECT EYE EXPOSURE

CLASS 3R LASER PRODUCT

LASER RADIATION

AVOID EXPOSURE TO BEAM

CLASS 3B LASER PRODUCT

LASER RADIATION

AVOID EYE OR SKIN EXPOSURE

TO DIRECT OR SCATTERED RADIATION

CLASS 4 LASER PRODUCT