# GUIDELINES ON THE USE OF GENERAL PURPOSE FUME CUPBOARDS

1.	Decide whether a general purpose fume cupboard is suitable for the proposed operation/experiment/reaction.  (Refer to COSHH risk assessment form).

2.	If it is suitable, turn on the fume cupboard fan and ensure the fume cupboard is operating satisfactorily.  (Read the airflow indicator, in the absence of this indicator, fit a “tell-tale” ribbon to the bottom of the sash.  This will provide a useful indication of airflow).

3.	Raise the sash to its maximum height.

4.	Bring all the equipment and chemicals required to the cupboard and place them in the working chamber of the fume cupboard.

5.	Avoid all unnecessary clutter in the fume cupboard.

6.	Do not block any exhaust slots or openings.

7.	Do not set up the apparatus/equipment close to the front edge of the fume cupboard.

(Ensure it is set back at least 150 mm from the plane of the sash).

8.	Do not position the equipment so far back that it obstructs the bottom “back baffle” slot or that the experimenter is required to put his/her head in the working chamber to operate the equipment.

9.	If flammable liquids are involved, use “flame-proof” electrical equipment.  Do not use multiple electric socket boards within the working chamber.

10.	Avoid as far as reasonably practicable naked flames (even when working with non-flammables), sources of high heat load, high-speed rotating machinery (eg centrifuges) and hot-air dryers.  (These items will disturb the air flow pattern and reduce the overall efficiency of containment of the fume cupboard.

11.	All large objects (safety screens, ovens, trays etc) should be raised about 50 mm off the work surface with blocks.  This will help to reduce turbulence.

12.	Reduce the sash opening to the recommended working height for the fume cupboard (typically 500 mm) and start the operation.

13.	Reduce the sash opening to the lowest practicable level as soon as possible.

14.	Whilst the experiment is in progress minimise movement of the sash.  Make any necessary movements of the sash slowly.

15.	Avoid rapid movements in the area adjacent to the fume cupboard.  Avoid leaning close to the front edge of the fume cupboard.

16.	Avoid rapid hand movements when making adjustments to the apparatus inside the fume cupboard.

17.	Plan for possible emergencies.  Fire in fume cupboards should be approached with extreme caution.  The use of high pressure CO2 extinguishers can result in flames being blown out of the fume cupboard (via the air make-up grille).  Materials can also be ejected in the direction of the operation.

18.	Avoid sitting at a fume cupboard.  This restricts mobility in the event of an emergency.

19.	Do not leave hazardous processes in the fume cupboard unattended.

20.	Clean up any spillages within the fume cupboard promptly and effectively.

21.	When the process is complete, keep the sash at the lowest level for a further 10 minutes to ensure that all fumes and vapours have been cleared before raising the sash to its maximum position.

- 2 -

22.	If applicable, allow the glassware/apparatus to cool.

23.	Dismantle the apparatus and set it aside for cleaning.

24.	Tidy up the fume cupboard.  Remove all equipment that is no longer needed.  Return chemicals to their proper storage place.

25.	Do not use the fume cupboard as a storage area.

26.	If appropriate, decontaminate and clean down the walls and worktop of the fume cupboard chamber.

27.	Lower the sash to its lowest working position.

28.	If appropriate, turn off the fume cupboard fan.